2018-6-20

# Guidelines for sequencing SCC libraries

All libraries made using SCC supplied hydrogel after 1-1-2017 are V3 libraries

Upon receiving your libraries:

When your libraries are delivered you will be given the library concentration as we determined using Qubit as well as the library index used for each sample. You will also be given the library size as determined by Bioanalyzer by and a Bioanalyzer trace showing a representative example of your libraries. You will NOT receive traces of all your libraries.

Each of your libraries will also have a separate library index.  We suggest pooling your samples by the concentration you are given then having the sequencing core do qPCR quantitation of the pooled sample for optimal cluster generation.

IF your libraries show any peaks below 200bp you will receive notification that your libraries likely have primer dimer contamination, which can negatively impact sequencing.  If this is significant we suggest you pool your samples at 2x or greater the concentration you will use for sequencing and do a final 0.8x SPRI cleanup on the sample (note there will be sample loss doing this).

This cleanup step can be done by the user, the SCC can help, or it is likely the sequencing core could perform this as part of QC.

Once you make your pool and are ready to sequence see the DETAILED instructions below as sequencing InDrops libraries:

For V3 sequencing please ask the core to deliver fastq files for ALL reads – including index reads. Or you can ask for the BCL files

V3 Sequencing

The V3 design aim was to get rid of the need for custom sequencing primers, include more diversity in certain parts of the barcode for higher quality sequencing, and save cycles by not sequencing the constant regions of the cell barcode.

V3 libraries have been successfully run on MiSeq, HiSeq and NextSeq instruments. While HiSeq usually gives the best data, this comes at a higher cost per base. These libraries are most often run on NextSeqs and the specific parameters for setting up a NextSeq run are given below:

PhiX is not necessary in sequencing.

For V3 sequencing on a NextSeq 75 cycle kit the reads are as follows:

Read 1 – reads into the mRNA

Read 2 – barcode/UMI to polyA tail

Index 1 – part of barcode

Index 2 – library index

NextSeq 75 cycle kit run of V3 libraries

please give your core the following set up parameters:

Platform   NextSeq

RE-HYB   No

Read type   Paired-end

Cycles read1   61

Cycles read2   14

Indexing   DualIndex

Cycles index 1   8

Cycles index 2   8

No custom primers

The i5 indexes (Index 2) are the library indexes we give to each sample or group of up to 3000 cells.

The cell barcode is made up of two 8pb sequences that are random combinations of 384 8bp sequences. The first of these is sequenced in the i7 sequencing read in our samples (single-cell barcodes), which are read as Index 1. The second half of the cell barcode is read in the first 8bp of the Read2 followed by the 6bp UMI read.

When setting up your sample sheet for the run you should just enter an arbitrary sequence for i5 and i7 (Index 1 and 2). Don’t worry about demultiplexing the sample and generating fastq files.  This demultiplexing will be done afterwards bioinformatically.

You either need the core to give you the fastq files for all reads – Read 1 and 2 and Index 1 and 2 – as well as reads from the undetermined file. If your core can not give you these files then you need the RAW BCL files for all four sequencing reads. What this means is that all of your data will go to one “undetermined file” off the instrument because none of the index reads will match what the instrument is expected to see. Pipelines for demultiplexing all of this exist through the Chan Bioinformatics core.

These are the sequences of the sample indices used on V3 libraries:

V3 Index 1 sequences

idx1	CTCTCTAT

idx2	TATCCTCT

idx 3	GTAAGGAG

idx 4	ACTGCATA

idx 5	AAGGAGTA

idx 6	CTAAGCCT

idx 7	CGTCTAAT

idx 8	TCTCTCCG

idx 9	TCGACTAG

idx 10	TTCTAGCT

idx 11	CCTAGAGT

idx 12	GCGTAAGA

idx 13	CTATTAAG

idx 14	AAGGCTAT

idx 15	GAGCCTTA

idx 16	TTATGCGA

idx 17   GGAGGTAA

idx 18   CATAACTG

idx 19   AGTAAAGG

idx 20   TAATCGTC

idx 21   TCCGTCTC

idx 22   AGCTTTCT

idx 23   AAGAGCGT

idx 24   AGAATGCG

InDrop Designs:  V2 vs V3

There are two InDrop designs used by the core.  As of September 2016 libraries will be made with the new V3 design.  Please be sure you know which version of the hydrogel was used on your sample as this makes VERY IMPORTANT DIFFERENCES IN SEQUENCING.

Libraries are prepared with custom primers that includes a 6-base (V2) or 8-base (V3) long library index. Read 1 adapter is on the 5’/cDNA end (reads gene), and Read 2 adapter is on the 3’ barcoded end (reads BC/UMI).

General requirements for all Illumina sequencers:

- Give your core facility the average library size (which will be given to you with your final libraries) and make sure the core facility performs library quantification qPCR before run.  It is fine to pool your libraries based on the concentrations given to you from the SCC with your library, but the final pool should be qPCR quantified before running on the sequencer. This will result in optimal cluster generation.

- V2 ONLY - Provide the sequencing core with custom primers (for Read 1, Index Read, and Read 2), or check to see if they already have them in stock. Check with the facility to see how much they need (the DFCI core requests at least 6 µL of 100 µM primer per run). See next page for primer sequences and index sequences.

Do NOT mix these custom primers into the standard Illumina primer mix.  This will cause sequencing Failure for InDrop libraries. V2 custom primers must be kept separate.

- V3 libraries require that fastq files be delivered for ALL reads – including index reads and undetermined reads.  If core facility cannot give you the fastq files for all 4 reads then ask for the bcl files.

V2 Sequencing

Our suggested method for sequencing is the Illumina NextSeq.  Users typically pool 10,000 to 30,000 cells on one NextSeq run.

Note that the Dana Farber Molecular Biology Core is well versed with InDrops libraries and has all custom primers required for V2 sequencing.  (https://mbcf.dfci.harvard.edu/?page_id=4)

Do NOT mix these custom primers into the standard Illumina primer mix.  This will cause sequencing Failure for InDrop libraries. V2 custom primers must be kept separate.

For NextSeq run:

- Use high-yield 75 cycle kit (which comes with 92 cycles)

- 36 cycles on read 1

- 6 cycles on index read

- Remaining 50 cycles on read 2

For diagnostic MiSeq run:

- Use v3 150 cycle kit (all kits comes with +15 cycles for index reads)

- Up to 103 cycles on read 1

- 6 cycles on index read

- 55 cycles on read 2

For HiSeq 2000 run:

- Samples must be multiplexed for this machine, so you are restricted to standard read lengths (be sure to include an index read)

- Use v3 reagents (v4 reagents may be incompatible with our library prep)

- Minimum required reads are Index, 37 cycles R1 and 50 cycles R2.  Read 1 can be longer if desired as this end reads into the gene

For HiSeq 2500 run:

- Same as MiSeq.

Sequencing primers (should be HPLC-purified):

Read_1_seq: GGCATTCCTGCTGAACCGCTCTTCCGATCT

Index_i7_seq: AGATCGGAAGAGCGTCGTGTAGGGAAAGAG

Read_2_seq: CTCTTTCCCTACACGACGCTCTTCCGATCT

Do NOT mix these custom primers into the standard Illumina primer mix.  This will cause sequencing Failure for InDrop libraries. V2 custom primers must be kept separate.

These are the sequences of the sample indices used on V2 libraries:

V2 Index sequences

idx1: ATCACG

idx2: CGATGT

idx3: TTAGGC

idx4: TGACCA

idx5: ACAGTG

idx6: GCCAAT

idx7: CAGATC

idx8: ACTTGA

idx9: GATCAG

idx10: TAGCTT

idx11: GGCTAC

idx12: CTTGTA

idx13: AGTCAA

idx14: AGTTCC

idx15: ATGTCA

idx16: CCGTCC

idx17: GTAGAG

idx18: GTCCGC

idx19: GTGAAA

idx20: GTGGCC

idx21: GTTTCG

idx22: CGTACG

idx23: GAGTGG

idx24: GGTAGC